Statement of Work (SOW)

Contract Title

Patient Elopement and Wander Management System (Replacement and Addition)

Background

The existing Patient Elopement and Wander Management System at the Dallas and Bonham Community Living Centers (CLC) require replacement due to age and technological deficiencies. This effort will also expand elopement monitoring coverage into select acute care areas within the Dallas VA Medical Center (5C400, 6C400, 7C200)

Scope

The Dallas VA Medical Center, part of the VA North Texas Healthcare System, requires the purchase and installation of a Patient Elopement and Wander Management System. The system will replace the existing system in the Dallas CLC (Building 60) and Bonham CLC (Building 29), and will expand into acute care areas 5C400, 6C400, and 7C200 in Dallas. The system will encompass the areas and specifics are detailed below.

Software and Networking Enterprise software must be virtualized onto a facility-supplied server located at the facility's data center, communicating via the Hospital's Local Area Network (LAN). Facility maps will be installed, and all system devices will be programmed into the system. The vendor will provide remote support access software for connectivity to technical support staff. The facility will be responsible for providing a working internet connection and IP address at the server location.

Client Workstations and Monitoring Displays Client workstations with proximity card readers for user login and touchscreen monitors must be installed and configured for enterprise use, communicating on the Hospital's LAN. A Staff Alert Panel (SAP) or similar may be installed at each coverage area as an analog backup for the computer system; in the event of a computer failure, the SAP will notify staff of any door and zone alarms. Workstations must be installed at each nurses' station within the coverage areas. All workstations must be networked to one another and to the server via an IT-provided VLAN switch.

Door and Zone Control Designated door zones must be capable of locking down using magnetic locks when a monitored patient is detected and the doors are closed. Not all doors require locking capability; the specific doors requiring lock integration will be identified during walkthroughs. Zone controllers must alarm and, where applicable, close and lock doors when a monitored patient enters the zone. All designated locked doors must remain locked until staff relocates the patient. Each door with locking capability may include a keypad or touchscreen display on both the protected and unprotected sides to clear alarms or grant emergency access via individual staff PIN codes. Virtual hallway door zones with motion control will also be monitored. Where conduit is required, procurement and installation of conduit will be the responsibility of the facility; the vendor will supply drawings to assist at no additional cost.

Alarm Identification and Reporting The system must store patient roaming history and zone status. Alarms must be displayed on a CAD drawing-based interface showing the patient's name, photo, and any associated warning messages. Upon alarm, the system must identify the zone or area where the patient triggered the alarm. The system must support customizable, web-based reports printable from a standard printer. The system must also provide GEO tracking capability for monitored patients.

Transmitters Patient wander transmitters with patient ID and/or banding must be supplied for use across all monitored units.

Integrations The system must integrate with the following:

- Existing fire alarm system (all designated locked doors must automatically unlock upon fire alarm activation). All fire release integration, permits, and associated costs are the responsibility of the facility.

- Existing security systems

Training The vendor must provide in-service instruction and go-live support for clinical personnel across all covered areas. The time and date for training will be established upon receipt of a purchase order.

Installation The vendor is responsible for all equipment installation and site preparation, including any electrical and data work required to make the system fully operational. Installation of all required networking equipment is included in this scope. The vendor will provide limited Dust Containment/Infection Control measures throughout the installation. Vendors will need to move ceiling tiles if necessary and vacuum the work areas upon completion.

- Infection Control Risk Assessment (ICRA) – All installation activities must comply with the facility's Infection Control Risk Assessment (ICRA) requirements. Prior to beginning work, the vendor must coordinate with the facility's Infection Control team to determine the applicable ICRA category for each work area and implement all required containment and protective measures accordingly. Work must not proceed in any area until ICRA requirements have been reviewed and approved by the facility's Infection Control representative. All installation activities must adhere to Class III dust containment requirements in accordance with the facility's Infection Control Risk Assessment (ICRA).

- Fire-Retardant Plywood – Where backboards are required for mounting head-end equipment, the vendor may reuse existing fire-retardant painted plywood provided it is structurally sound and in good condition. Reuse of existing plywood is permitted as a means of maintaining the room's existing fire rating. Any plywood determined to be damaged, deteriorated, or otherwise unsuitable for reuse must be replaced with new plywood that meets the same fire-retardant rating as the original material. The vendor must document the condition of any reused plywood and make that documentation available to the COR upon request.

Equipment Specifics

Equipment needed for these areas will depend on coverage required by the site. Included in this requirement are floor maps of the areas where the installations will be performed.

- Patient Wrist Bands – Quantity dependent on rooms identified as high-risk areas

- Server Bay Installation: Install rack mountable 1U server in provided 19” rack space and connect to provided ports to switch.

- Nursing Staff Training: Training for the Nursing Staff will be provided for each shift.

- Old Equipment Removal: When new equipment has been installed and tested, the old equipment can be removed and turned over to Nursing Staff for proper disposal.

- Deliverables: Deliverables will be received by the warehouse and delivered to a designated space until supplies are needed for installation.

- All head-end equipment and computer systems provided under this contract must be connected to a vendor-supplied and installed Uninterruptible Power Supply (UPS) of sufficient capacity to maintain system operation during a loss of primary power.

Performance Monitoring

- The COR or designated VA Engineering Employee will visually inspect daily for quality of work and adherence to the SOW, ILSM, and ICRA.

- VA Staff will verify daily cleanliness of the area before, after, and during the workday.

Other Pertinent Information or Special Considerations

- All penetrations through walls will be filled with VA approved fire caulk. All penetrations will be inspected by VA Safety Personnel.

- A copy of the approved Interim Life Safety Measures, (ILSM) and Infections Control Risk Assessment, (ICRA) documents must be always displayed during the installation.

- Work shall be accomplished Monday-Friday 8:00 am - 5:00 pm.

National Holidays: The 10 holidays observed by the Federal Government are:

  - New Year Day

  - Martin Luther King's Birthday

  - Presidents Day

  - Memorial Day

  - Independence Day

  - Labor Day

  - Columbus Day

  - Veterans Day

  - Thanksgiving Day

  - Christmas Day

  - And any other day specifically declared by the President of the United States to be a national holiday

- Contractor must check in with the Biomedical Engineering Service in the Bldg. 2J BB220 each day prior to starting work.

- System must be installed without disturbing the existing system. If existing system must be disturbed, a three-business-day lead time will be required to bring on extra staff to cover the entrance/exit.

- A Site Visit is mandatory before a bid will be accepted.

- Inspection and Acceptance Criteria.

  - Nursing Staff training for all shifts will be accepted by the COR, after the Chief of Community Living Center and Clinical Leadership in Bldg. 2 6C and 7C-400 are satisfied that the staff on different shift are trained.

  - Biomedical Staff training will be accepted by the COR, after the Biomedical Supervisor is satisfied there was acceptable training.

  - The COR or designated VA Staff will verify correct operations, programming, and training before installation will be complete.

- The Patient Elopement System must be delivered on a platform capable of supporting future expansion into asset and equipment tracking and staff duress functionality

Risk Control

- When pulling wire in areas where patients will be present, minimize the amount of cabling on the ground at one time and utilize safety cones when there is a trip hazard due to cabling.

- Pulling of wire where it may make a potential trip hazard will only be performed in slack traffic periods.

- Full compliance with the ICRA and ILSM shall be upheld.

Hours of Performance

Delivery and services shall occur Monday through Friday 8:00 AM to 5:00 PM CST for all project tasks, unless otherwise scheduled with Biomedical Engineering.

Period of Performance

Period of Performance:  August 1, 2026 – July 31, 2027 (One year)

Dates as specified unless contract award passes the start date. After which, the start date will be the day of award.

Check in Requirements

The Field Service Engineer must report to the Biomedical Engineering Department to obtain a badge and sign in with the Biomedical Engineering Service before work begins.  Submit any mobile media devices that would be used on the system to a virus scan.  Upon completion of work, the Field Service Engineer must report to the Biomedical Engineering Service to sign out and brief Biomedical Staff or Supervisor if Biomedical Staff is unavailable concerning completion of service.  At the end of briefing Field Service Engineer will return the badge and sign out in the Biomedical Engineering Department.

Documentation

Contractor shall furnish a detailed field service report upon completion of work to Biomedical Staff in the Biomed Department. Payment will not be processed until a properly completed service report is received.  The service report shall contain, at a minimum, the following information:

- Type, model, System ID, and serial number (s) of all equipment on which maintenance was performed

- Total time spent performing maintenance.

- Detailed narrative description of the services required

- Copies of all test reports

- Complete list of parts replaced (when applicable)

- Date and time the repair/support was completed

The service report shall itemize every item in the specification.  Each item shall state the "as found" condition or values, the "calibrated to" or "adjusted to" values, the factory design tolerances, and a complete description of all work performed concerning the items.  Included will be a list of new parts used and recommended future repairs.

Place of Performance

Department of Veterans Affairs

Dallas VA Medical Center

4500 S. Lancaster Road

Dallas, Texas 75216

Contracting Officer Representatives

Prior to contract award, the Contracting Officer shall designate a VA Medical Center employee as the COR.  All work coordination shall be made through the COR.  The Contractor shall be provided a copy of the letter of delegation authorizing the COR at the commencement of the term of the contract.  No other person shall be authorized to act in such capacity unless appointed in writing by the Contracting Officer.

Information Technology Security Requirements

The contractor, their personnel, and their subcontractors shall be subject to the Federal laws, regulations, standards, and VA Directives and Handbooks regarding information and information system security as delineated in this contract. The contractor shall comply with all Federal laws and regulations the VA has developed when VA sensitive information is accessed, used, stored, generated, transmitted, or exchanged by and between VA and a contractor. The information made available to the contractor by VA for the performance of this contract will be used only for the purposes of performance under this contract.   The certification and accreditation requirements do not apply to this requirement and a security accreditation package is not required.

Security Statement

Contractor will check in with the Biomedical Engineering Service prior to beginning work each day.

The Dallas VA Medical Center Police Service will determine the status of each Contractor or Affiliate’s background investigation status for the issuance of a Non-PIV or PIV-Flash card.

Contractor will not have access to PHI or PII information or sources.

Contractor will be escorted at all times, by Authorized VA Engineering Staff, while they are performing work in the IDF or Server Bay.

Sensitive VA information is contained within the systems covered by this contract.

Biomedical Engineering shall perform virus scans on all removable media prior to use on VA medical equipment. This includes all types of removable media, including media (e.g., USB devices, CDs, dongles, etc.) that has been issued by VA, media not issued by VA, and media brought in by vendors or independent service organizations. Within accordance of VA Directive 6500, Information Security Program, September 2007

The Vendor will not transfer any VA information to a location outside the VA and only to VA locations determined by the VA System Administrator.  The information in these systems may be covered by the Privacy Act 1974 which contains criminal penalties of abuse of information.

During onsite service, the Vendor shall be chaperoned by VA Personnel.  However, the vendor shall not be issued a User ID/Password.

Non-volatile memory devices, working or non-working, shall NOT be removed from the VA Medical Center Dallas until the ISO has certified that the data has been destroyed.   For magnetic devices and media, the data destruction will be by degaussing.  Other forms of cleansing will be used for non-magnetic media.

The vendor will not have remote access to complete the repair(s) and preventive maintenance.

- GENERAL

Contractors, contractor personnel, subcontractors, and subcontractor personnel shall be

subject to the same Federal laws, regulations, standards, and VA Directives and Handbooks

as VA and VA personnel regarding information and information system security.

2. ACCESS TO VA INFORMATION AND VA INFORMATION SYSTEMS

a. A contractor/subcontractor shall request logical (technical) or physical access to VA

information and VA information systems for their employees, subcontractors, and affiliates only

to the extent necessary to perform the services specified in the contract, agreement, or task

order.

b. All contractors, subcontractors, and third-party servicers and associates working with

VA information are subject to the same investigative requirements as those of VA appointees

or employees who have access to the same types of information. The level and process of

background security investigations for contractors must be in accordance with VA Directive

and Handbook 0710, Personnel Suitability and Security Program. The Office for Operations,

Security, and Preparedness is responsible for these policies and procedures.

c. Contract personnel who require access to national security programs must have a valid

security clearance. National Industrial Security Program (NISP) was established by Executive

Order 12829 to ensure that cleared U.S. defense industry contract personnel safeguard the

classified information in their possession while performing work on contracts, programs, bids,

or research and development efforts. The Department of Veterans Affairs does not have a

Memorandum of Agreement with Defense Security Service (DSS). Verification of a Security

Clearance must be processed through the Special Security Officer located in the Planning and

National Security Service within the Office of Operations, Security, and Preparedness.

d. Custom software development and outsourced operations must be located in the U.S.

to the maximum extent practical. If such services are proposed to be performed abroad and

are not disallowed by other VA policy or mandates, the contractor/subcontractor must state

where all non-U.S. services are provided and detail a security plan, deemed to be acceptable

by VA, specifically to address mitigation of the resulting problems of communication, control,

data protection, and so forth. Location within the U.S. may be an evaluation factor.

e. The contractor or subcontractor must notify the Contracting Officer immediately when

an employee working on a VA system or with access to VA information is reassigned or leaves

the contractor or subcontractor’s employ. The Contracting Officer must also be notified

immediately by the contractor or subcontractor prior to an unfriendly termination.

3. VA INFORMATION CUSTODIAL LANGUAGE

a. Information made available to the contractor or subcontractor by VA for the performance

or administration of this contract or information developed by the contractor/subcontractor in

performance or administration of the contract shall be used only for those purposes and shall

not be used in any other way without the prior written agreement of the VA. This clause

expressly limits the contractor/subcontractor's rights to use data as described in Rights in Data

- General, FAR 52.227-14(d) (1).

b. VA information should not be co-mingled, if possible, with any other data on the

contractors/subcontractor’s information systems or media storage systems in order to ensure

VA requirements related to data protection and media sanitization can be met. If co-mingling

must be allowed to meet the requirements of the business need, the contractor must ensure

that VA’s information is returned to the VA or destroyed in accordance with VA’s sanitization

requirements. VA reserves the right to conduct on site inspections of contractor and

subcontractor IT resources to ensure data security controls, separation of data and job duties,

and destruction/media sanitization procedures are in compliance with VA directive

requirements.

c. Prior to termination or completion of this contract, contractor/subcontractor must not

destroy information received from VA, or gathered/created by the contractor in the course of

performing this contract without prior written approval by the VA. Any data destruction done on

behalf of VA by a contractor/subcontractor must be done in accordance with National Archives

and Records Administration (NARA) requirements as outlined in VA Directive 6300, Records

and Information Management and its Handbook 6300.1 Records Management Procedures,

applicable VA Records Control Schedules, and VA Handbook 6500.1, Electronic Media

Sanitization. Self-certification by the contractor that the data destruction requirements above

have been met must be sent to the VA Contracting Officer within 30 days of termination of the

contract.

d. The contractor/subcontractor must receive, gather, store, back up, maintain, use,

disclose and dispose of VA information only in compliance with the terms of the contract and

applicable Federal and VA information confidentiality and security laws, regulations and

policies. If Federal or VA information confidentiality and security laws, regulations and policies

become applicable to the VA information or information systems after execution of the

contract, or if NIST issues or updates applicable FIPS or Special Publications (SP) after

execution of this contract, the parties agree to negotiate in good faith to implement the

information confidentiality and security laws, regulations and policies in this contract.

e. The contractor/subcontractor shall not make copies of VA information except as

authorized and necessary to perform the terms of the agreement or to preserve electronic

information stored on contractor/subcontractor electronic storage media for restoration in case

any electronic equipment or data used by the contractor/subcontractor needs to be restored to

an operating state. If copies are made for restoration purposes, after the restoration is

complete, the copies must be appropriately destroyed.

f. If VA determines that the contractor has violated any of the information confidentiality,

privacy, and security provisions of the contract, it shall be sufficient grounds for VA to withhold payment to the

contractor or third party or terminate the contract for default or terminate for

cause under Federal Acquisition Regulation (FAR) part 12.

g. If a VHA contract is terminated for cause, the associated BAA must also be terminated

and appropriate actions taken in accordance with VHA Handbook 1600.01, Business

Associate Agreements. Absent an agreement to use or disclose protected health information,

there is no business associate relationship.

h. The contractor/subcontractor must store, transport, or transmit VA sensitive information

in an encrypted form, using VA-approved encryption tools that are, at a minimum, FIPS 140-2

validated.

i. The contractor/subcontractor’s firewall and Web services security controls, if applicable,

shall meet or exceed VA’s minimum requirements. VA Configuration Guidelines are available

upon request.

j. Except for uses and disclosures of VA information authorized by this contract for

performance of the contract, the contractor/subcontractor may use and disclose VA information

only in two other situations: (i) in response to a qualifying order of a court of competent

jurisdiction, or (ii) with VA’s prior written approval. The contractor/subcontractor must refer all

requests for, demands for production of, or inquiries about, VA information and information

systems to the VA contracting officer for response.

k. Notwithstanding the provision above, the contractor/subcontractor shall not release VA

records protected by Title 38 U.S.C. 5705, confidentiality of medical quality assurance records

and/or Title 38 U.S.C. 7332, confidentiality of certain health records pertaining to drug

addiction, sickle cell anemia, alcoholism or alcohol abuse, or infection with human

immunodeficiency virus. If the contractor/subcontractor is in receipt of a court order or other

requests for the above-mentioned information, that contractor/subcontractor shall immediately

refer such court orders or other requests to the VA contracting officer for response.

l. For service that involves the storage, generating, transmitting, or exchanging of VA

sensitive information but does not require C&A or an MOU-ISA for system interconnection, the

contractor/subcontractor must complete a Contractor Security Control Assessment (CSCA) on

a yearly basis and provide it to the COTR.

GENERAL RULES OF BEHAVIOR

a. Rules of Behavior are part of a comprehensive program to provide complete

information security. These rules establish standards of behavior in recognition of the fact that

knowledgeable users are the foundation of a successful security program. Users must

understand that taking personal responsibility for the security of their computer and the

information it contains is an essential part of their job.

b. The following rules apply to all VA contractors. I agree to:

(1) Follow established procedures for requesting, accessing, and closing user accounts

and access. I will not request or obtain access beyond what is normally granted to users or

by what is outlined in the contract.

(2) Use only systems, software, databases, and data which I am authorized to use,

including any copyright restrictions.

(3) I will not use other equipment (OE) (non-contractor owned) for the storage, transfer,

or processing of VA sensitive information without a VA CIO approved waiver, unless it has

been reviewed and approved by local management and is included in the language of the

contract. If authorized to use OE IT equipment, I must ensure that the system meets all

applicable 6500 Handbook requirements for OE.

(4) Not use my position of trust and access rights to exploit system controls or access

information for any reason other than in the performance of the contract.

(5) Not attempt to override or disable security, technical, or management controls

unless expressly permitted to do so as an explicit requirement under the contract or at the

direction of the COTR or ISO. If I am allowed or required to have a local administrator

account on a government-owned computer, that local administrative account does not confer

me unrestricted access or use, nor the authority to bypass security or other controls

except as expressly permitted by the VA CIO or CIO's designee.

(6) Contractors’ use of systems, information, or sites is strictly limited to fulfill the terms

of the contract. I understand no personal use is authorized. I will only use other Federal

government information systems as expressly authorized by the terms of those systems. I

accept that the restrictions under ethics regulations and criminal law still apply.

(7) Grant access to systems and information only to those who have an official need to

know.

(8) Protect passwords from access by other individuals.

(9) Create and change passwords in accordance with VA Handbook 6500 on systems

and any devices protecting VA information as well as the rules of behavior and

security settings for the particular system in question.

(10) Protect information and systems from unauthorized disclosure, use, modification, or

destruction. I will only use encryption that is FIPS 140-2 validated to safeguard VA sensitive

information, both safeguarding VA sensitive information in storage and in transit regarding

my access to and use of any information assets or resources associated with my

performance of services under the contract terms with the VA.

(11) Follow VA Handbook 6500.1, Electronic Media Sanitization to protect VA

information. I will contact the COTR for policies and guidance on complying with this

requirement and will follow the COTR's orders.

(12) Ensure that the COTR has previously approved VA information for public

dissemination, including e-mail communications outside of the VA as appropriate. I will not

make any unauthorized disclosure of any VA sensitive information through the use of any

means of communication including but not limited to e-mail, instant messaging, online chat,

and web bulletin boards or logs.

(13) Not host, set up, administer, or run an Internet server related to my access to and

use of any information assets or resources associated with my performance of services

under the contract terms with the VA unless explicitly authorized under the contract or in

writing by the COTR.

(14) Protect government property from theft, destruction, or misuse. I will follow VA

directives and handbooks on handling Federal government IT equipment, information, and

systems. I will not take VA sensitive information from the workplace without authorization

from the COTR.

(15) Only use anti-virus software, antispyware, and firewall/intrusion detection software

authorized by VA. I will contact the COTR for policies and guidance on complying with this

requirement and will follow the COTR's orders regarding my access to and use of any

information assets or resources associated with my performance of services under the

contract terms with VA.

(16) Not disable or degrade the standard anti-virus software, antispyware, and/or

firewall/intrusion detection software on the computer I use to access and use information

assets or resources associated with my performance of services under the contract terms

with VA. I will report anti-virus, antispyware, firewall or intrusion detection software errors, or

significant alert messages to the COTR.

(17) Understand that restoration of service of any VA system is a concern of all users of

the system.

(18) Complete required information security and privacy training, and complete required

training for the particular systems to which I require access.

PRIVACY

- Contractors and any subcontractors must adhere to the provisions of Public Law 104-191, Health Insurance Portability and Accountability Act (HIPAA) of 1996.  This includes both the Privacy and Security Rules published by the Department of Health and Human Services (HHS).

- As required by HIPAA, HHS has promulgated rules governing the use and disclosure of protected health information by covered entities, Veterans Health Administration (VHA). In accordance with HIPAA, the contractor may be required to enter into a Business Associate Agreement (BAA) with VHA.

- Business associates must follow VHA privacy policies and practices when applicable. All contractors and business associates must receive privacy training annually.

- For contractors and business associates who do not have access to VHA computer systems, this requirement is met by completing VHA National Privacy Policy training, other VHA approved privacy training or contractor furnished training that meets the requirements of the HHS Standards for Privacy of Individually Identifiable Health Information as determined by VHA.

- For contractors and business associates who are granted access to VHA computer systems, this requirement is met by completing VHA National Privacy Policy training or other VHA approved privacy training.  Proof of training is required upon request.

Records Management Statement

1.  Contractor shall comply with all applicable records management laws and regulations, as well as National Archives and Records Administration (NARA) records policies, including but not limited to the Federal Records Act (44 U.S.C. chs. 21, 29, 31, 33), NARA regulations at 36 CFR Chapter XII Subchapter B, and those policies associated with the safeguarding of records covered by the Privacy Act of 1974 (5 U.S.C. 552a). These policies include the preservation of all records, regardless of form or characteristics, mode of transmission, or state of completion.

2.  In accordance with 36 CFR 1222.32, all data created for Government use and delivered to, or falling under the legal control of, the Government are Federal records subject to the provisions of 44 U.S.C. chapters 21, 29, 31, and 33, the Freedom of Information Act (FOIA) (5 U.S.C. 552), as amended, and the Privacy Act of 1974 (5 U.S.C. 552a), as amended and must be managed and scheduled for disposition only as permitted by statute or regulation.

3.  In accordance with 36 CFR 1222.32, Contractor shall maintain all records created for Government use or created in the course of performing the contract and/or delivered to, or under the legal control of the Government and must be managed in accordance with Federal law. Electronic records and associated metadata must be accompanied by sufficient technical documentation to permit understanding and use of the records and data.

4. VA Medical Center Dallas and its contractors are responsible for preventing the alienation or unauthorized destruction of records, including all forms of mutilation. Records may not be removed from the legal custody of  the VA Medical Center Dallas or destroyed except for in accordance with the provisions of the agency records schedules and with the written concurrence of the Head of the Contracting Activity. Willful and unlawful destruction, damage or alienation of Federal records is subject to the fines and penalties imposed by 18 U.S.C. 2701.

In the event of any unlawful or accidental removal, defacing, alteration, or destruction of records, Contractor must report to VA Medical Center Dallas. The agency must report promptly to NARA in accordance with 36 CFR 1230.

5.  The Contractor shall immediately notify the appropriate Contracting Officer upon discovery of any inadvertent or unauthorized disclosures of information, data, documentary materials, records or equipment. Disclosure of non-public information is limited to authorized personnel with a need-to-know as described in the [contract vehicle]. The Contractor shall ensure that the appropriate personnel, administrative, technical, and physical safeguards are established to ensure the security and confidentiality of this information, data, documentary material, records and/or equipment is properly protected. The Contractor shall not remove material from Government facilities or systems, or facilities or systems operated or maintained on the Government’s behalf, without the express written permission of the Head of the Contracting Activity. When information, data, documentary material, records and/or equipment is no longer required, it shall be returned to VA Medical Center Dallas control or the Contractor must hold it until otherwise directed. Items returned to the Government shall be hand carried, mailed, emailed, or securely electronically transmitted to the Contracting Officer or address prescribed in the [contract vehicle]. Destruction of records is EXPRESSLY PROHIBITED unless in accordance with Paragraph (4).

6.  The Contractor is required to obtain the Contracting Officer's approval prior to engaging in any contractual relationship (sub-contractor) in support of this contract requiring the disclosure of information, documentary material and/or records generated under, or relating to, contracts. The Contractor (and any sub-contractor) is required to abide by Government and VA Medical Center Dallas guidance for protecting sensitive, proprietary information, classified, and controlled unclassified information.

7.  The Contractor shall only use Government IT equipment for purposes specifically tied to or authorized by the contract and in accordance with VA Medical Center Dallas policy.

8.  The Contractor shall not create or maintain any records containing any non-public VA Medical Center Dallas information that are not specifically tied to or authorized by the contract.

9.  The Contractor shall not retain, use, sell, or disseminate copies of any deliverable that contains information covered by the Privacy Act of 1974 or that which is generally protected from public disclosure by an exemption to the Freedom of Information Act.

10.  The VA Medical Center Dallas owns the rights to all data and records produced as part of this contract. All deliverables under the contract are the property of the U.S. Government for which VA Medical Center Dallas shall have unlimited rights to use, dispose of, or disclose such data contained therein as it determines to be in the public interest. Any Contractor rights in the data or deliverables must be identified as required by FAR 52.227-11 through FAR 52.227-20.

11.  Training.  All Contractor employees assigned to this contract who create, work with, or otherwise handle records are required to take VHA-provided records management training, Talent Management System (TMS) Item #3873736, Records Management for Records Officers and Liaisons. The Contractor is responsible for confirming training has been completed according to agency policies, including initial training and any annual or refresher training.

NARA Records Management Language for Contracts (May 2017)

1.  Contractor shall comply with all applicable records management laws and regulations, as well as National Archives and Records Administration (NARA) records policies, including but not limited to the Federal Records Act (44 U.S.C. chs. 21, 29, 31, 33), NARA regulations at 36 CFR Chapter XII Subchapter B, and those policies associated with the safeguarding of records covered by the Privacy Act of 1974 (5 U.S.C. 552a). These policies include the preservation of all records, regardless of form or characteristics, mode of transmission, or state of completion.

2.  In accordance with 36 CFR 1222.32, all data created for Government use and delivered to, or falling under the legal control of, the Government are Federal records subject to the provisions of 44 U.S.C. chapters 21, 29, 31, and 33, the Freedom of Information Act (FOIA) (5 U.S.C. 552), as amended, and the Privacy Act of 1974 (5 U.S.C. 552a), as amended and must be managed and scheduled for disposition only as permitted by statute or regulation.

3.  In accordance with 36 CFR 1222.32, Contractor shall maintain all records created for Government use or created in the course of performing the contract and/or delivered to, or under the legal control of the Government and must be managed in accordance with Federal law. Electronic records and associated metadata must be accompanied by sufficient technical documentation to permit understanding and use of the records and data.

4. VA North Texas Health Care System and its contractors are responsible for preventing the alienation or unauthorized destruction of records, including all forms of mutilation. Records may not be removed from the legal custody of VA North Texas Health Care System or destroyed except for in accordance with the provisions of the agency records schedules and with the written concurrence of the Head of the Contracting Activity. Willful and unlawful destruction, damage or alienation of Federal records is subject to the fines and penalties imposed by 18 U.S.C. 2701.

In the event of any unlawful or accidental removal, defacing, alteration, or destruction of records, Contractor must report to VA North Texas Health Care System.  The agency must report promptly to NARA in accordance with 36 CFR 1230.

5.  The Contractor shall immediately notify the appropriate Contracting Officer upon discovery of any inadvertent or unauthorized disclosures of information, data, documentary materials, records or equipment. Disclosure of non-public information is limited to authorized personnel with a need-to-know as described in the [contract vehicle]. The Contractor shall ensure that the appropriate personnel, administrative, technical, and physical safeguards are established to ensure the security and confidentiality of this information, data, documentary material, records and/or equipment is properly protected. The Contractor shall not remove material from Government facilities or systems, or facilities or systems operated or maintained on the Government’s behalf, without the express written permission of the Head of the Contracting Activity. When information, data, documentary material, records and/or equipment is no longer required, it shall be returned to VA North Texas Health Care System control or the Contractor must hold it until otherwise directed. Items returned to the Government shall be hand carried, mailed, emailed, or securely electronically transmitted to the Contracting Officer or address prescribed in the [contract vehicle]. Destruction of records is EXPRESSLY PROHIBITED unless in accordance with Paragraph (4).

6.  The Contractor is required to obtain the Contracting Officer's approval prior to engaging in any contractual relationship (sub-contractor) in support of this contract requiring the disclosure of information, documentary material and/or records generated under, or relating to, contracts. The Contractor (and any sub-contractor) is required to abide by Government and VA North Texas Health Care System guidance for protecting sensitive, proprietary information, classified, and controlled unclassified information.

7.  The Contractor shall only use Government IT equipment for purposes specifically tied to or authorized by the contract and in accordance with VA North Texas Health Care System policy.

8.  The Contractor shall not create or maintain any records containing any non-public VA North Texas Health Care System information that are not specifically tied to or authorized by the contract.

9.  The Contractor shall not retain, use, sell, or disseminate copies of any deliverable that contains information covered by the Privacy Act of 1974 or that which is generally protected from public disclosure by an exemption to the Freedom of Information Act.

10.  The VA North Texas Health Care System owns the rights to all data and records produced as part of this contract. All deliverables under the contract are the property of the U.S. Government for which VA North Texas Health Care System shall have unlimited rights to use, dispose of, or disclose such data contained therein as it determines to be in the public interest. Any Contractor rights in the data or deliverables must be identified as required by FAR 52.227-11 through FAR 52.227-20.

11.  Training.  All Contractor employees assigned to this contract who create, work with, or otherwise handle records are required to take VHA-provided records management training, Talent Management System (TMS) Item #10176, Privacy and Information Security, Rules of Behavior. The Contractor is responsible for confirming training has been completed according to agency policies, including initial training and any annual or refresher training.

Floor Plans