SPECIFICATIONS

- Specifications Description:

Provide a preparative scale High-Performance Liquid Chromatography System for natural products discovery for laboratory use. Ideally the system would be modular with the ability to upgrade specific options in the future, such as the type of detector. It is anticipated that the instrument will not be connected to the internet or local area network.

Minimum requirements for the HPLC system we require:

- The system should have all essential parts of an HPLC, to include: pump(s), a UV-Vis detector (200-800 nm), an Evaporative Light Scattering Detector, a controller, and all the necessary tubing and accessories required to run the system, including the software necessary to connect to the controller and run the system. All major parts of the system (I.E. the pumps, detector, controller, etc.) should be from the same manufacturer or at least supported (as in troubleshooting support) by the same manufacturer.

- The system should include a fraction collector with the ability to accommodate a variety of standard test tube sizes and specifically 13 x 100 mm and 16 x 150 mm, with a minimum of 48 fractions

- The system must have multiple pumps (at least 2), or at absolute minimum a single pump capable of generating gradients from multiple inputs (at least 2). They should support flow rates from 1.0 ml/min to 150 ml/min. Minimum pressure tolerance of at least 5000 psi.

- Specific non-detector potions of the system:

  - The solvent delivery system should be able to operate in constant flow or constant pressure modes.

  - The pumps must have a high flow rate accuracy (+ or - 2%).

  - The pumps should come equipped with a suction filter to prevent particulate contamination.

- The HPLC will require 2 specific types of detector,a  UV-Vis detector with a range of 200 -800 nm and an Evaporative Light Scattering Detector. The detectors should either be manufactured by the same manufacturer as the other major parts or at least be supported (as in Troubleshooting support) by the same manufacturer as the other major parts. The system is preferred to be modular and allows the installation of additional detectors in future.

- The instrument should have a column selector with ability to select and switch between at least three different columns.

- Any HPLC option considered should include technical installation and initial testing of the system. The installation should include basic training on the software used to control the system. The technician performing the installation is required to be a U.S. Citizen.

- Additional System Specifications:

- System Controller

- Solvent Delivery Pump

  - Auto-diagnostics to detect problems during analysis, and an auto-recovery function

- Electrical requirements are North American 115 V 60hz

- There should be a warranty of at least 1 year, and preferably 2 years or more, for the system.

- The vendor will provide a technician for installation and training of NCAUR personnel.

2.0 Place of Delivery

USDA/ARS

1815 N. University Street

Peoria IL 61604

United States

3.0 Delivery Schedule

| Item | QTY | Unit | Place of Delivery | Delivery Date |
| --- | --- | --- | --- | --- |
| HPLC System | 1 |  | 1815 N University | 7/30/2026 |

4.0 Delivery Point of Contact

Chris Dunlap, Heather Walker

5.0 Additional Requirements

Technician provided for installation will need verification.