STATEMENT OF WORK

Title: Oscilloscope

- BACKGROUND INFORMATION

The semiconductor supply chain is global, specialized, and interconnected. Chipmakers do business with thousands of individual suppliers that provide the highly complex materials and tools used to produce semiconductors. To address the lack of full visibility into the semiconductors markets supply chain and R&D ecosystem gaps the National Institute for Standards and Technology (NIST) will conduct the measurement science, or metrology, critical to the development of new materials, packaging, and production methods in chip manufacturing.

The National Institute of Standards and Technology (NIST) Materials Science and Engineering Division (MSED) within the Material Measurement Laboratory (MML) is involved in the development of standard reference materials (SRMs) for magnetic moment – a critical standard for measuring magnetic materials for next generation memory and computing. The SRMs are certified using the NIST-built Absolute Magnetometer, a unique custom instrument. For SI-traceability, the force acting on the SRMs will be measured and then traced back to the kilogram. To measure the force accurately and repeatably, this work requires an oscilloscope. The oscilloscope will be integrated into NIST’s Absolute Magnetometer and must communicate with the control computer of the Absolute Magnetometer system. The development of the Absolute Magnetometer for oscillatory fields will enable the certification of new magnetic SRMs that can be used by other government agencies (i.e. FDA, EPA in regulations), industry, and academia to calibrate their instrumentation to make quantitative and comparable measurements for faster development of new technologies (i.e., next generation magnetic memory and computing, as well as other fields including medical imaging and diagnostics and non-destructive evaluation of materials for real-time monitoring of infrastructure).

- PURPOSE

The contractor shall provide one (1) oscilloscope inclusive of FOB destination delivery and minimum 1-year warranty.

- MINIMUM REQUIREMENTS

The Contractor shall provide a system that meets all technical specifications identified below. All items must be new. Used or remanufactured equipment will not be considered for award. Experimental, prototype, or custom items will not be considered.  The use of “gray market” components are not acceptable. All line items shall be shipped in the original manufacturer’s packaging and include all original documentation and software, when applicable.

Line Item 0001: Oscilloscope

Quantity: One (1)

- Technical Specifications

  - Bandwidth (operating frequency range): up to at least 2.5 GHz

  - Sampling Rate: at least 20 GSa/s (giga-samples per second) on two (analog or digital) channels

  - Record length: at least 1 Gpts/channel (giga-points per channel) on two channels

  - ADC Resolution: 12 bits (Please specify ENOB)

  - Vertical Resolution: at least 8 bits at maximum bandwidth and sample rate

  - Input: at least four analog channels

  - Input Impedance:

    - 1 MΩ ± 1%

    - 50 Ω ± 3% (or less)

  - Input Capacitance: approx. 14 pF

  - Input Sensitivity:

    - 1 MΩ: at least 1 mV/division to 10 V/division

    - 50 Ω: 1 mV/division to 1 V/division

  - Jitter: ≤ 80 fs

  - Time Accuracy: ±1.0x10-7

  - Computer Connectors: USB

  - Power: 115 V/60 Hz/single phase; cable must be provided

  - Probes (including any associated connectors or amplifiers):

    - Passive X10 probes with a bandwidth of at least 0.5 GHz (quantity = 4)

    - High current (range of at least 5 mA - 100 A, DC and AC up to at least 10 MHz) Probe (quantity = 4)

  - Software capable of downloading data V(t) or I(t)

  - Calibration traceable to the SI

  - Two (2) sets of manuals.  The manual may be supplied in hard copy or electronic format (or both). The manual must cover the proper operation, routine maintenance, and troubleshooting for the instrument, all components, and controlling software

Line Item 0002: X10 Probes

Quantity: Four (4)

- Technical Specifications

  - Bandwidth: ≥0.5 GHz

  - Attenuation Factor: 10X

  - Input Capacitance: <11 pF

  - Cable Length: ≥3 m

  - Includes all connectors to work with oscilloscope

Line Item 0003: High Current Probes

Quantity: Four (4)

- Technical Specifications

  - Bandwidth: DC to ≥5 MHz

  - Maximum Current: 150 ARMS

  - Peak Pulse Current: 500 A

  - Cable Length: ≥3 m

  - Includes all amplifiers and connectors to work with oscilloscope

- SCHEDULE OF DELIVERABLES

| Deliverable Number | Description | Quantity | Due Date | Place of Delivery |
| --- | --- | --- | --- | --- |
| 1 | Oscilloscope with power cable | One (1) | No later than 6 months after receipt of order. | NIST Gaithersburg |
| 2 | X10 Probes | Four (4) | No later than 6 months after receipt of order. | NIST Gaithersburg |
| 3 | High Current Probes | Four (4) | No later than 6 months after receipt of order. | NIST Gaithersburg |
| 4 | Manuals | Two (2) | No later than 6 months after receipt of order. | NIST Gaithersburg |

PLACE OF PERFORMANCE

All work shall be completed at the Contractor’s facility

PERIOD OF PERFORMANCE

The delivery date shall be six (6) months from award of this requirement.

- DELIVERY TERMS

Delivery shall be F.O.B Destination (or equivalent incoterms, such as DDP) and shall occur in accordance with the delivery due dates provided in the above table.

FOB Destination means:  The contractor shall pack and mark the shipment in conformance with carrier requirements, deliver the shipment in good order and condition to the point of delivery specified in the purchase order, be responsible for any loss of and/or damage to the goods occurring before receipt and acceptance of the shipment by the consignee at the delivery point specified in the purchase order; and pay all charges to the specified point of delivery.

The contractor shall deliver all Line Items to:

National Institute of Standards and Technology

Shipping and Receiving

100 Bureau Drive, Building 301

Gaithersburg, MD 20899

C/O Cindi Dennis 223/B144

- INSPECTION & ACCEPTANCE

In addition to the inspection and acceptance terms articulated in the specific FAR clause 52.212-4, the Government reserves the right to perform such performance tests and evaluations to verify specified system performance. Such tests and evaluations, if performed, shall be conducted within the environment that the system is to be operated. The Contractor has the right to be present during the tests and evaluations, if performed, at the Contractor’s expense. NIST may choose at its discretion to forego this part of acceptance testing.

A visual inspection of the equipment will be performed by the NIST Technical Point of Contract (TPOC) to identify surface defects or any form of indication that any equipment was damaged during transport to NIST. The Government shall have sole discretion to require repair or replacement of damaged and/or nonconforming supplies at no cost to the Government. The Government at any time prior to acceptance shall reject the equipment due to defects and/or nonconformance. The vendor is responsible for latent defects discovered any time after final inspection. However, the extent of its liability shall be prorated over the useful life of the equipment.

Ownership of the equipment shall transfer to NIST upon acceptance by the Government.

The Government will test, inspect, and accept or reject the equipment within 4 weeks of the receipt of the equipment unless otherwise indicated above. The Government reserves the right to conduct quality assurance testing to confirm that a given instrument(s) meets the manufacturer’s and/or the Government’s performance specifications. It is anticipated that the equipment will meet all manufacturer’s specifications and/or the Government’s performance specifications identified in the most recent operations and maintenance manual for each piece of equipment and/or in this document.

- WARRANTY

The contractor shall warranty the entire system for a period of a minimum of one (1) year after receipt of the equipment and shall be in accordance with terms in FAR 52.212-4. Warranty shall commence upon acceptance of the system by the Government and at a minimum shall include all materials, labor, and shipping costs.

- PAYMENT SCHEDULE

The Contractor will be paid, in accordance with the payments clause in the contract and as otherwise noted in this document, upon receipt of a proper invoice and after successful demonstration by the instrumentation that it performs in accordance with the requirements set forth in this document.